Stat-Pak Competency Assessment Test (CAT) Checklist

Competency Assessment testing is required for the Stat-Pak HIV Rapid Test initially on site after certification training, at 6 months and annually thereafter to maintain certification.  The CAT must be directly observed and signed by the Authorized Evaluator.  Completed CAT checklists must be filed in the technicians personnel file

and a copy must be faxed to 415-934-4868.

Date:				   Site Name and Number:

Test Reader’s Name: 				   		            Tech ID:                             .

	Gathers/arranges all materials

- Examines test kit pouch (unopened, room temperature, absorbent packet)

- Accurately records lot number of test kit on supplemental lab slip

	Accurately records expiration date of test kit on supplemental lab slip

	Accurately records lot number of buffer solution on supplemental lab slip

	Accurately records expiration date of buffer solution on supplemental lab slip

- Accurately records Tech ID on supplemental lab slip

- Opens the test kit packet and places it on an absorbent pad (chuck) to run the test

- Places the OraQuick lab slip sticker in the appropriate place on the supplemental lab slip

- Places the OraQuick lab slip sticker on the Stat-Pak test kit

- Wears gloves for all subsequent steps

- Wears lab coat and face mask for the four following steps:

- Opens the top of the vacutainer tube or the control vial, without spilling contents

- Carefully inserts loop into the tube or control vial

- Verifies that the loop is completely filled and that no bubbles exist in the loop

- Recaps the vacutainer tube or control vial

- Touches the tip of the loop to the sample well in the Stat-Pak test kit, until all sample is dispensed

- Adds 3 free-flowing drops of Stat-Pak buffer, holding the buffer container vertically

	Accurately records start time and temperature on lab slip (temperature range within 64 – 80 ºF)

- Verifies that the ID number on the lab slip and the test kit match

- Accurately records end time and temperature on lab slip  (time range within 15-20 minutes)

- Accurately records result

	Checks lab slip for completeness and accuracy of information including test result

	Successfully completes all steps (if not, note what was missing/incorrect below):

# Text boxes

□  Initial CAT	□ 6 month CAT	  □ Annual CAT									7/31/2009

Notes:

Authorized Evaluator Name:

(Print Full Name)

Evaluator’s Signature:

____Initial CAT         ____6 month CAT         ____Annual CAT

□  Initial CAT	□ 6 month CAT	  □ Annual CAT									7/31/2009